Thursday Feb. 3, 2011

I would encourage you to throw away the notes from class last Thursday.  Here's what I hope is a clearer and more understandable explanation of ion counters and conductivity meters.

Ion counters measure the concentration of small ion charge carriers in the air (as we will see they can only measure one polarity of charge carrier at a time).  A conductivity meter measures the air's conductivity.  And, actually, it isn't able to measure the total conductivity because both positively and negatively charged small ions make the air conductive.  A conductivity meter is only able to measure the contribution that small ions of a single polarity make.  

Ions counters and conductivity meters often make used of the cylindrical capacitor geometry shown below.

We sort of implicity assume that we are looking a just a small section in the middle of a much longer cylinder.  The E field has just a radial component in this case.  The solution we obtain above is the same as we got for the field around an infinitely long line of charge with uniform line charge density.

We can write the electric field in terms of the potential difference between the two cylinders.

If the outer cylinder is connected to positive potential, the grounded inner cylinder will have negative charged induced on it.  The E field will point inward as shown in the figure above.  It is a relatively easy matter to determine the capacitance

We'll look first at how an ion counter works.

The electric field will cause a  positively charged small ion entering the cylindrical capacitor at the left will drift toward the inner conductor with a drift velocity (vd).  In time dt, the small ion will drift dr (Point 1).  The drift velocity, Point 2, is just the electrical mobility, Be, times the electric field.  T at Point 3 is the time it will take the small ion to drift from the outer cylinder to the inner cylinder.

Now in order to be "counted" the small ion must make it to the center conductor before it travels a distance L (refer back to the first figure in today's notes).

This means (Point 4) that T must be less than L/u, where u is the speed at which the air is traveling along the length of the cylindrical capacitor (we assume u is uniform, that there is no dependence on r).

At Point 5 we can see that, for a given value of u, Be must be greater than a certain critical mobility, Bc, in order to be collected (i.e. make it to the center conductor before traveling the length of the cylindrical capacitor).

We can write Bc in terms of a volumetric flow rate (which might actually be easier to measure than u).  This is shown at Point 6.

We want Bc to be small so that all of the small ions have a mobility greater than Bc and can be counted.  Clearly the lower the flow rate and the longer the tube, the more time the small ions will spend in the capacitor and the more likely they will be collected.  Increasing the potential difference between the two cylinders will increase the strength of the electric field and the inward drift velocity of the small ions.

Here's the relationship between a measured signal current (and known volumetric flow rate) and the small ion concentration (actually, you are able to determine the concentration times the charge per small ion).

When functioning as a conducitivty meter, only the small ions in a portion of the volume of air flowing through the cylindrical capacitor are collected (the green shaded volume in the figure below).

The expression at Point 7 is really just an earlier expression (Point 3 on one of the earlier figures) with rc replacing b/2.

At Point 8 we are multiplying the rate at which the green shaded volume is flowing through the cylinder times N q.  This is the charge collected at the inner conductor per unit time and is the signal current.  We need to try to relate the signal current to conductivity.

We substitute in for capacitance and recognize that N Be q is conductivity and obtain the expressionshown at Point 9.  There is a linear relationship between isignal and V.  The slope of a plot of isignal vs V should provide an estimate of conductivity.  And again because we collect only one polarity of small ion we aren't measuring the total conductivity.  The total conductivity depends on charge carriers of both polarities.

The following figure shows a conventional ion counter/conductivity meter design and an op-amp circuit that could be used to measure the signal current (Vout  = signal current x R).

Note how the op-amp keeps the center conductor at ground potential.  Because the signal current is very small, a large feedback resistor is needed in the op-amp circuit (1013 ohms was used in the "instrument" brought to class).

In this last figure we get a better idea of what the transition from conductivity meter to ion counter means.  

For a given rate of air flow through the cylindrical capacitor we monitor the signal current as the potential difference between the outer and inner conductors is increased.  As V increases small ions in a growing volume of air are collected and measured.  The signal current increases.  Eventually all of the small ions are collected and the signal current flattens out.

The slope of the linearly increasing, early portion of the plot provides an estimate of conductivity.  The amplitude of the signal current, once it has flattened out, can be related to small ion concentration.